SPECIГО

LIS Integration Engine

Foundation Document

| Analyser | Beckman Coulter AU480 |
| --- | --- |
| Lab | MMI Diagnostics, Raipur |
| Interface Direction | Unidirectional (Analyser -> LIMS) |
| Stack | Node.js / MySQL / React.js |
| OS at Site | Windows (on-site PC at MMI) |
| Doc Version | v1.5 ¦ April 2026 - FINAL. LIMS integration points, UI spec, multi-manufacturer framework added. |

# 1. Purpose of This Document

This is the master reference document for the SpeciGo LIS Integration Engine, to be uploaded into the Claude Project workspace. It captures all confirmed information about the Beckman Coulter AU480 analyser at MMI Diagnostics, the SpeciGo LIMS data model, the confirmed communication settings from the analyser screen, and the complete build plan. Every future conversation within the Claude Project should reference this document.

This document must be updated after each integration milestone is completed. Keep it current so every new Claude session starts with accurate context.

# 2. Analyser Master Record - Beckman Coulter AU480

This is the complete canonical record for the AU480. These values populate the analyser_master table in MySQL.

## 2.1 Identity Block

| Field | Value |
| --- | --- |
| Manufacturer | Beckman Coulter |
| Model Name | AU480 |
| Model Family | AU Series (shared spec with AU680) |
| Category | Clinical Chemistry Analyser |
| Analyser Type | Photometric + ISE (Potentiometry) |
| Throughput | 400 photometric tests/hr ¦ 800 with ISE |
| Simultaneous Tests | Up to 63 analytes onboard |
| Sample Types | Serum, Urine, Whole Blood, Other (X/Y) |
| Serial Number | 2023070728 |
| Firmware Version | [To be confirmed at site] |
| Manufacture Date | 2023-07-06 (July 2023) ¦ Made in Japan |
| ISE Module Present | [To be confirmed at MMI - Na, K, Cl] |
| Site - Lab Name | MMI Diagnostics |
| Site - City | Raipur, Chhattisgarh |

## 2.2 Communication Block - CONFIRMED FROM MMI MACHINE PHOTOS

All values below are confirmed directly from the AU480 screen photos taken at MMI Diagnostics on 08 April 2026. These are not defaults or assumptions.

| Parameter | Value (from MMI screen) | Source Screen |
| --- | --- | --- |
| Interface Type | RS-232C (Serial) + LAN both connected | Photo - Protocol tab |
| Protocol Standard | Beckman Coulter proprietary (ASTM-based) | AU480 Online Spec v9 |
| Communication Class | Class B (with ACK/NAK) | Photo - Protocol tab |
| Baud Rate | 9600 bps | Photo - Protocol tab |
| Character Length | 8 bits | Photo - Protocol tab |
| Parity Bit | No (None) | Photo - Protocol tab |
| Stop Bit | 1 bit | Photo - Protocol tab |
| Start Code (1st) | 02h (STX) | Photo - Protocol tab |
| End Code (1st) | 03h (ETX) | Photo - Protocol tab |
| Max Text Length | 512 bytes | Photo - Protocol tab |
| BCC Check | Disabled (unchecked) | Photo - Protocol tab |
| ETB Control | Disabled (unchecked) | Photo - Protocol tab |
| Retry Count | 3 | Photo - Protocol tab |
| T.R.I. Receive Error | Continue | Photo - Protocol tab |
| Results Transfer Error | Continue | Photo - Protocol tab |

## 2.3 Format Configuration - CONFIRMED FROM MMI MACHINE PHOTOS

| Parameter | Value (from MMI screen) | Notes |
| --- | --- | --- |
| Rack No./Cup Position | ENABLED (checked) | Format Config tab |
| Type field | ENABLED (checked) | Format Config tab |
| Dilution Info | DISABLED (unchecked) | Format Config tab |
| Reagent Info | DISABLED (unchecked) | Format Config tab |
| R1-2/R2-2 Use | DISABLED (unchecked) | Format Config tab |
| Zero Suppress | ENABLED (checked) | Format Config tab |
| Rack No. Digit | 4 digits | Format Config tab |
| Online Test No. Digit | 3 digits (001-120) | Format Config tab |
| Result Digit | 6 digits | Format Config tab |
| No. of Data Marks | 2 types | Format Config tab |
| Cal/Control No. Digit | 3 digits | Format Config tab |

## 2.4 Sample ID Configuration - CONFIRMED

| Parameter | Value (from MMI screen) | Notes |
| --- | --- | --- |
| Sample ID Mode | Barcode | Analysis Mode screen - S.ID Barcode section |
| Barcode Type | MULTI CODE | Analysis Mode screen |
| Check Mode | No (No Chk.Chr.) | Analysis Mode screen |
| Test Requisition Mode | Sequential (Routine, Emergency, STAT) | Analysis Mode screen |
| Default Sample Type | Serum | Analysis Mode screen - Others section |
| STAT Operation | Manual | Analysis Mode screen |
| Auto Repeat - Rack | Disabled | Analysis Mode screen |
| Auto Repeat - STAT | Disabled | Analysis Mode screen |

IMPORTANT: Both RS-232 serial and LAN cables are physically connected at MMI. The machine may be configured to use one or both. We need to confirm which the lab intends to use for our integration. For our initial integration, we will target RS-232 serial as it is simpler and confirmed active. LAN (TCP/IP) integration is the fallback path.

## 2.5 Online Test Number to Parameter Mapping - FROM MMI MACHINE

The following are the ACTIVE online test numbers configured at MMI Diagnostics, extracted from the Online Test No. screen photo taken on-site. Only populated entries are shown.

| Online No. | Test Name | Online No. | Test Name |
| --- | --- | --- | --- |
| 001 | W340 (Wavelength 340) | 002 | W380 |
| 003 | W480 | 004 | W600 |
| 005 | R150 | 006 | S1.0ul |
| 007 | S2.0ul | 008 | S5.0ul |
| 009 | ALB (Albumin) | 010 | ALP |
| 011 | ALT | 012 | AMY (Amylase) |
| 013 | AST | 014 | CAL (Calcium) |
| 015 | CHOL (Cholesterol) | 016 | CKN (CK-NAC) |
| 017 | CRE (Creatinine) | 018 | DBB (Direct Bilirubin) |
| 019 | DBC | 020 | GGT |
| 021 | GLU (Glucose) | 022 | HDL |
| 023 | PHO (Phosphorus) | 024 | FE (Iron/Ferritin) |
| 025 | MG (Magnesium) | 026 | TBB (Total Bilirubin) |
| 027 | TBC | 028 | TP (Total Protein) |
| 029 | TRG (Triglycerides) | 030 | UREA |
| 031 | UA (Uric Acid) | 032 | ULB |
| 033 | GLB (Globulin) | 034 | UBC |
| 096 | LIH | 097 | Na (Sodium) |
| 098 | K (Potassium) | 099 | CI (Chloride) |

The Online Test No. is what appears in the message data field (11) in all analysis data messages. Our parser must use this 3-digit number to look up the correct LIMS parameter. Numbers 035-095 and 100-120 are not populated at MMI and will produce empty or absent data in messages.

# 3. AU480 Message Format - Technical Reference

All analysis data from the AU480 uses the Beckman Coulter proprietary online protocol documented in the AU680/AU480 Online Specification v9 (Jan 2011). This is ASTM-inspired but NOT standard ASTM E1394. Do not use a generic ASTM parser without understanding these differences.

## 3.1 Message Structure

Every message from the AU480 follows this structure:

[STX]  [2-char distinction code]  [system no]  [message data]  [ETX]

      
STX = 0x02 (start byte)
ETX = 0x03 (end byte)
BCC = disabled at MMI (no checksum byte appended)
Class B = AU480 expects ACK (0x06) response after each block
         NAK (0x15) signals retransmission request

## 3.2 Message Flow for a Normal Sample (Unidirectional Mode)

In our unidirectional setup, the AU480 pushes data to our system. The sequence for a complete sample result is:

| Step | Code | Meaning |
| --- | --- | --- |
| 1 | DB | Analysis data transmission start - AU480 signals it is about to send results |
| 2 | D\u0394 or DH | Normal sample data OR Repeat run data - one per sample, contains all results |
| 3 | d\u0394 | STAT quick output (if STAT samples run) |
| 4 | DR | Reagent blank data (if configured) |
| 5 | DA | Calibration data (if configured) |
| 6 | DQ | QC data (if QC run) |
| 7 | DE | Analysis data transmission end - session closed |

For our unidirectional integration, we only care about D-type messages (DB, D space, DH, DE). The R-type messages (sample information request/response) are for bidirectional mode only. We do not send S-type responses.

## 3.3 Analysis Data Message Structure - What Our Parser Reads

The critical message is the D-space (normal sample data) message. It has a fixed part and a variable part:

| Field | Size | Description |
| --- | --- | --- |
| Message Start Code | 1 byte | 0x02 (STX) - confirmed at MMI |
| Message Distinction Code | 2 chars | D + space for normal sample data |
| System No. | 0 or 2 chars | Device identifier (disabled at MMI) |
| Rack No. | 4 digits | Rack position e.g. 0001. Confirmed 4-digit setting. |
| Cup Position | 2 digits | 01-10 for rack samples, 01-22 for STAT |
| Sample Type | 1 char | Space=Serum, U=Urine, X=Other, W=Whole blood |
| Sample No. | 4 chars | 0001-9999 Routine, E001-E999 Emergency, P001-P999 STAT |
| Sample ID | Configurable | The barcode value scanned at sample intake |
| Dummy | 4 chars | Spaces (0x20 x4) - always present |
| Data Classification No. | 1 char | 0-9 for blocks, E for last block |
| [Variable Part starts here] |  |  |
| Sex | 1 char | M/F/0/space - part of variable data |
| Year Age | 3 chars | 000-150 or spaces |
| Month Age | 2 chars | 00-11 or spaces |
| Patient Info 1-6 | Variable | Up to 20 chars each, if enabled |
| Online Test No. | 3 digits | 001-120 - identifies which parameter follows |
| Reagent Info (if enabled) | Lot + Bottle | Not enabled at MMI |
| Analysis Data | 6 digits | Numeric result value (Zero Suppress ON at MMI) |
| Data Flag | 2 chars | Error/status flag - 2 types at MMI |
| Message End Code | 1 byte | 0x03 (ETX) |

## 3.4 Result Value Format

Results are 6 digits with Zero Suppress enabled at MMI. This means leading zeros are replaced with spaces (0x20). For example:

Raw value: 045.25  -> transmitted as: ' 45.25'  (leading zero replaced by space)
Raw value: 000.00  -> transmitted as: '   .  '  (all zeros suppressed)
Overflow:  999999  -> analyser was above measurement range
Negative:  -12.34  -> stored as 7-char in 9-digit mode (not used at MMI - 6 digit mode)

## 3.5 Data Flag Reference

The 2-type data flag setting at MMI means each result has a 2-character flag field. The first character is the primary flag, second is secondary. Key flags to handle in our parser:

| Flag | Meaning |
| --- | --- |
| H | Result higher than reference range |
| L | Result lower than reference range |
| ph | Result higher than PANIC (critical high) value |
| pl | Result lower than PANIC (critical low) value |
| # | Insufficient sample volume |
| % | Clot detected in sample |
| ? | Unable to calculate result |
| F/G | Result above/below dynamic range (dilution required) |
| R | Insufficient reagent during run |
| N | Normal (within reference range) |
|  | Two spaces = no flag / not analyzed |

# 4. SpeciGo LIMS Database Schema Analysis

The following is derived from analysis of the SpeciGo LIMS MySQL database (u151751738_theranizeDevh1). This section maps the LIMS data model to where analyser results must land.

## 4.1 Existing Analyser Tables

The LIMS already has lab_analyzers and lab_analyzer_parameters tables. These exist but are minimal. We will EXTEND these rather than create new tables.

-- EXISTING: lab_analyzers
CREATE TABLE `lab_analyzers` (
  `id` bigint UNSIGNED NOT NULL,
  `analyzer_uid` varchar(20) NOT NULL,
  `analyzer_name` varchar(150) NOT NULL,
  `analyzer_code` varchar(100) NOT NULL,
  `ip` varchar(45) NOT NULL,         -- currently only IP stored
  `port` int(11) NOT NULL,            -- currently only port stored
  `lab_uid` varchar(50) NOT NULL,
  `is_active` tinyint(1) NOT NULL DEFAULT 1
)

      
-- EXISTING: lab_analyzer_parameters
CREATE TABLE `lab_analyzer_parameters` (
  `id` bigint UNSIGNED NOT NULL,
  `lab_uid` varchar(50) NOT NULL,
  `analyzer_code` varchar(100) NOT NULL,
  `parameter_code` varchar(100) NOT NULL,
  `lims_parameter_uid` varchar(50) NOT NULL,
  `lims_test_uid` varchar(50) NOT NULL
)

## 4.2 Required Schema Extensions

The existing lab_analyzers table is missing critical fields for serial communication, protocol details, and integration metadata. The following ALTER statements extend it:

-- EXTEND lab_analyzers with communication fields
ALTER TABLE `lab_analyzers`
  ADD COLUMN `manufacturer` varchar(100) DEFAULT NULL AFTER `analyzer_name`,
  ADD COLUMN `model` varchar(100) DEFAULT NULL AFTER `manufacturer`,
  ADD COLUMN `category` varchar(100) DEFAULT NULL AFTER `model`,
  ADD COLUMN `connection_type` enum('SERIAL','TCP_IP','BOTH') DEFAULT 'SERIAL' AFTER `category`,
  ADD COLUMN `serial_port` varchar(20) DEFAULT NULL AFTER `connection_type`,
  ADD COLUMN `baud_rate` int(11) DEFAULT 9600 AFTER `serial_port`,
  ADD COLUMN `data_bits` int(11) DEFAULT 8 AFTER `baud_rate`,
  ADD COLUMN `stop_bits` int(11) DEFAULT 1 AFTER `data_bits`,
  ADD COLUMN `parity` enum('NONE','EVEN','ODD') DEFAULT 'NONE' AFTER `stop_bits`,
  ADD COLUMN `flow_control` enum('NONE','HARDWARE','SOFTWARE') DEFAULT 'NONE' AFTER `parity`,
  ADD COLUMN `protocol` varchar(50) DEFAULT 'AU480' AFTER `flow_control`,
  ADD COLUMN `communication_class` enum('A','B') DEFAULT 'B' AFTER `protocol`,
  ADD COLUMN `bcc_check` tinyint(1) DEFAULT 0 AFTER `communication_class`,
  ADD COLUMN `start_code` varchar(10) DEFAULT '02h' AFTER `bcc_check`,
  ADD COLUMN `end_code` varchar(10) DEFAULT '03h' AFTER `start_code`,
  ADD COLUMN `text_length` int(11) DEFAULT 512 AFTER `end_code`,
  ADD COLUMN `online_test_no_digits` int(11) DEFAULT 3 AFTER `text_length`,
  ADD COLUMN `result_digits` int(11) DEFAULT 6 AFTER `online_test_no_digits`,
  ADD COLUMN `rack_no_digits` int(11) DEFAULT 4 AFTER `result_digits`,
  ADD COLUMN `interface_direction` enum('UNIDIRECTIONAL','BIDIRECTIONAL') DEFAULT 'UNIDIRECTIONAL' AFTER `rack_no_digits`,
  ADD COLUMN `integration_status` enum('NOT_STARTED','IN_PROGRESS','TESTED','LIVE') DEFAULT 'NOT_STARTED' AFTER `interface_direction`,
  ADD COLUMN `serial_number` varchar(100) DEFAULT NULL AFTER `integration_status`,
  ADD COLUMN `firmware_version` varchar(50) DEFAULT NULL AFTER `serial_number`,
  ADD COLUMN `notes` text DEFAULT NULL AFTER `firmware_version`;

-- EXTEND lab_analyzer_parameters with online test no mapping
ALTER TABLE `lab_analyzer_parameters`
  ADD COLUMN `online_test_no` varchar(10) DEFAULT NULL AFTER `parameter_code`,
  ADD COLUMN `unit` varchar(50) DEFAULT NULL AFTER `online_test_no`,
  ADD COLUMN `is_active` tinyint(1) DEFAULT 1 AFTER `unit`;

-- NEW TABLE: lis_integration_results
-- Raw results from analyser BEFORE LIMS mapping
CREATE TABLE `lis_integration_results` (
  `id` bigint UNSIGNED NOT NULL AUTO_INCREMENT,
  `lab_uid` varchar(50) NOT NULL,
  `analyzer_uid` varchar(20) NOT NULL,
  `sample_id` varchar(100) NOT NULL COMMENT 'barcode scanned at analyser',
  `sample_no` varchar(20) DEFAULT NULL COMMENT 'internal analyser sample no',
  `sample_type` varchar(5) DEFAULT NULL COMMENT 'S=Serum U=Urine X=Other',
  `rack_no` varchar(10) DEFAULT NULL,
  `cup_position` varchar(5) DEFAULT NULL,
  `online_test_no` varchar(10) NOT NULL COMMENT '3-digit test code from analyser',
  `parameter_code` varchar(100) DEFAULT NULL COMMENT 'mapped analyser param code',
  `lims_parameter_uid` varchar(50) DEFAULT NULL COMMENT 'resolved LIMS param UID',
  `raw_value` varchar(20) DEFAULT NULL COMMENT 'value as received from analyser',
  `numeric_value` decimal(12,4) DEFAULT NULL COMMENT 'parsed numeric value',
  `data_flag` varchar(10) DEFAULT NULL COMMENT 'flags: H L ph pl etc',
  `mapping_status` enum('MAPPED','UNMAPPED','ERROR') DEFAULT 'UNMAPPED',
  `lims_write_status` enum('PENDING','WRITTEN','FAILED','SKIPPED') DEFAULT 'PENDING',
  `lims_barcode_uid` varchar(50) DEFAULT NULL COMMENT 'matched report_barcode.barcode_uid',
  `result_category` enum('PATIENT','QC','CALIBRATION','REAGENT_BLANK') DEFAULT 'PATIENT',
  `control_id` varchar(50) DEFAULT NULL COMMENT 'for DQ messages - control barcode',
  `calibrator_id` varchar(50) DEFAULT NULL COMMENT 'for DA messages - calibrator ID',
  `control_no` varchar(10) DEFAULT NULL,
  `calibrator_no` varchar(10) DEFAULT NULL,
  `error_message` text DEFAULT NULL,
  `raw_message` text DEFAULT NULL COMMENT 'full raw analyser message for debugging',
  `received_at` timestamp NOT NULL DEFAULT current_timestamp(),
  `written_at` timestamp NULL DEFAULT NULL,
  PRIMARY KEY (`id`),
  INDEX `idx_sample_id` (`sample_id`),
  INDEX `idx_lab_analyzer` (`lab_uid`, `analyzer_uid`),
  INDEX `idx_lims_write_status` (`lims_write_status`)
) ENGINE=InnoDB DEFAULT CHARSET=utf8mb4;

-- NEW TABLE: lis_integration_log
-- Communication session and error log
CREATE TABLE `lis_integration_log` (
  `id` bigint UNSIGNED NOT NULL AUTO_INCREMENT,
  `lab_uid` varchar(50) NOT NULL,
  `analyzer_uid` varchar(20) NOT NULL,
  `session_type` enum('DB','DE','ERROR','ACK','NAK','CONNECT','DISCONNECT') NOT NULL,
  `message_code` varchar(10) DEFAULT NULL COMMENT 'e.g. D , DH, DB, DE',
  `sample_id` varchar(100) DEFAULT NULL,
  `details` text DEFAULT NULL,
  `created_at` timestamp NOT NULL DEFAULT current_timestamp(),
  PRIMARY KEY (`id`),
  INDEX `idx_lab_analyzer_time` (`lab_uid`, `analyzer_uid`, `created_at`)
) ENGINE=InnoDB DEFAULT CHARSET=utf8mb4;

## 4.3 How Results Flow Into Existing LIMS Tables

The matching logic when a result arrives from the AU480:

| Step | Action | LIMS Table |
| --- | --- | --- |
| 1. Receive | AU480 sends result with Sample ID (barcode) | lis_integration_results (raw) |
| 2. Match Barcode | SELECT * FROM report_barcode WHERE barcode_uid = sample_id (exact match - barcode_uid IS the Sample ID) | report_barcode |
| 3. Match Test | Use online_test_no to find lims_parameter_uid via lab_analyzer_parameters | lab_analyzer_parameters |
| 4. Write Value | UPDATE report_test_values_data SET value = numeric_value WHERE barcode_uid AND parameter_uid | report_test_values_data |
| 5. Update Status | UPDATE report_tests SET test_status = 'TC' (test completed) WHERE barcode_uid | report_tests |
| 6. Log | Insert into lis_integration_log and update lis_integration_results.lims_write_status | lis_integration_log |

CONFIRMED: SpeciGo prints the barcode_uid directly on the tube label (e.g. 69D5DCF585189 - 13 alphanumeric characters). The AU480 scans this and sends it as the Sample ID in every result message. Our lookup query is: SELECT * FROM report_barcode WHERE barcode_uid = [sample_id_from_analyser]. The AU480 Sample ID digit count must be set to 13 digits on the analyser.

# 5. Node.js Integration Engine - Build Plan

## 5.1 Project Folder Structure

speciго-lis-engine/
├── config/
│   └── analysers/
│       └── au480_mmi.json         <- AU480 instance config for MMI
├── src/
│   ├── transport/
│   │   ├── SerialPortManager.js   <- RS-232 connection handler
│   │   └── TCPSocketManager.js    <- LAN TCP connection handler (future)
│   ├── protocol/
│   │   ├── AU480Parser.js         <- Parses AU480 proprietary messages
│   │   └── MessageFramer.js       <- STX/ETX framing + ACK/NAK handling
│   ├── mapping/
│   │   └── ParameterMapper.js     <- Online Test No. -> LIMS parameter
│   ├── db/
│   │   ├── ResultWriter.js        <- Writes to report_test_values_data
│   │   └── IntegrationLogger.js   <- Writes to lis_integration_log
│   ├── api/
│   │   └── IntegrationAPI.js      <- REST endpoints for React dashboard
│   └── engine/
│       └── IntegrationEngine.js   <- Main orchestrator
├── test/
│   └── simulate_au480.js          <- Test script to simulate AU480 messages
└── index.js                       <- Entry point

## 5.2 AU480 Instance Config File - au480_mmi.json

This JSON config file is what Umesh fills in per installation. Only connection fields change between labs.

{
  "analyser_id": "AU480",
  "lab_uid": "[MMI_LAB_UID]",
  "analyzer_uid": "[FROM_lab_analyzers_TABLE]",
  "manufacturer": "Beckman Coulter",
  "model": "AU480",
  "site": "MMI Diagnostics, Raipur",
  "connection": {
    "type": "serial",
    "port": "COM3",
    "baudRate": 9600,
    "dataBits": 8,
    "stopBits": 1,
    "parity": "none",
    "flowControl": "none"
  },
  "protocol": {
    "class": "B",
    "startCode": "0x02",
    "endCode": "0x03",
    "bccCheck": false,
    "textLength": 512,
    "onlineTestNoDigits": 3,
    "resultDigits": 6,
    "rackNoDigits": 4,
    "dataMarks": 2
  },
  "parameter_map": {
    "009": { "code": "ALB",  "lims_name": "Albumin",       "unit": "g/dL" },
    "010": { "code": "ALP",  "lims_name": "ALP",           "unit": "U/L"  },
    "011": { "code": "ALT",  "lims_name": "ALT (SGPT)",    "unit": "U/L"  },
    "012": { "code": "AMY",  "lims_name": "Amylase",       "unit": "U/L"  },
    "013": { "code": "AST",  "lims_name": "AST (SGOT)",    "unit": "U/L"  },
    "014": { "code": "CAL",  "lims_name": "Calcium",       "unit": "mg/dL"},
    "015": { "code": "CHOL", "lims_name": "Total Cholesterol", "unit": "mg/dL"},
    "016": { "code": "CKN",  "lims_name": "CK-NAC",        "unit": "U/L"  },
    "017": { "code": "CRE",  "lims_name": "Creatinine",    "unit": "mg/dL"},
    "018": { "code": "DBB",  "lims_name": "Direct Bilirubin", "unit": "mg/dL"},
    "020": { "code": "GGT",  "lims_name": "GGT",           "unit": "U/L"  },
    "021": { "code": "GLU",  "lims_name": "Glucose",       "unit": "mg/dL"},
    "022": { "code": "HDL",  "lims_name": "HDL Cholesterol", "unit": "mg/dL"},
    "023": { "code": "PHO",  "lims_name": "Phosphorus",    "unit": "mg/dL"},
    "024": { "code": "FE",   "lims_name": "Iron",          "unit": "ug/dL"},
    "025": { "code": "MG",   "lims_name": "Magnesium",     "unit": "mg/dL"},
    "026": { "code": "TBB",  "lims_name": "Total Bilirubin", "unit": "mg/dL"},
    "028": { "code": "TP",   "lims_name": "Total Protein", "unit": "g/dL" },
    "029": { "code": "TRG",  "lims_name": "Triglycerides", "unit": "mg/dL"},
    "030": { "code": "UREA", "lims_name": "Urea",          "unit": "mg/dL"},
    "031": { "code": "UA",   "lims_name": "Uric Acid",     "unit": "mg/dL"},
    "033": { "code": "GLB",  "lims_name": "Globulin",      "unit": "g/dL" },
    "097": { "code": "NA",   "lims_name": "Sodium",        "unit": "mEq/L"},
    "098": { "code": "K",    "lims_name": "Potassium",     "unit": "mEq/L"},
    "099": { "code": "CL",   "lims_name": "Chloride",      "unit": "mEq/L"}
  },
  "message_filters": {
    "patient_results": true,
    "stat_results": true,
    "qc_results": false,
    "calibration_results": false,
    "reagent_blank": false
  }
}

# 6. Complete Build Sequence

Claude will generate every file listed below in full. Umesh monitors and deploys. Each step is self-contained with complete code provided.

| Step | Module | Task | Status |
| --- | --- | --- | --- |
| 1 | Database | Run ALTER statements to extend lab_analyzers and lab_analyzer_parameters | Pending |
| 2 | Database | Create lis_integration_results and lis_integration_log tables | Pending |
| 3 | Config | Create au480_mmi.json with confirmed settings from this doc | Pending |
| 4 | Transport | Build SerialPortManager.js using 'serialport' npm package | Pending |
| 5 | Protocol | Build MessageFramer.js - STX/ETX detection + ACK/NAK response | Pending |
| 6 | Protocol | Build AU480Parser.js - parse DB, D-space, DE messages | Pending |
| 7 | Mapping | Build ParameterMapper.js - online_test_no -> lims_parameter_uid | Pending |
| 8 | DB Write | Build ResultWriter.js - write to lis_integration_results, then LIMS tables | Pending |
| 9 | Test | Build simulate_au480.js - send fake AU480 messages over serial/TCP | Pending |
| 10 | API | Build GET /api/lis/results and GET /api/lis/status endpoints | Pending |
| 11 | Frontend | Build React LIS Dashboard component | Pending |
| 12 | Site Test | Test at MMI with actual AU480 connection | Pending |

## 6.1 Required npm Packages

npm install serialport          # RS-232 serial port communication
npm install @serialport/parser-byte-length  # for framing
npm install mysql2              # MySQL database driver
npm install express             # REST API
npm install dotenv              # environment config
npm install winston             # logging

# 7. Open Items and Decisions Required

ALL 7 ITEMS RESOLVED. Item 4 is deferred to go-live as a data configuration step via the admin UI. All remaining items are confirmed. Build can begin immediately.

| # | Question | Why It Matters | Status |
| --- | --- | --- | --- |
| 1 | What COM port does the RS-232 cable use on the Windows PC at MMI? | CONFIRMED: COM2 | DONE |
| 2 | What barcode format does SpeciGo print on sample tube labels? | CONFIRMED: barcode_uid directly - 13 chars e.g. 69D5DCF585189 | DONE |
| 3 | Is LAN intended as primary or just backup at MMI? | CONFIRMED: Serial (RS-232 on COM2) is primary. LAN is available but not used for integration. | DONE |
| 4 | What is the lims_parameter_uid for each active test at MMI? | DEFERRED TO GO-LIVE: This is data configuration not code. Done via admin UI mapping screen at site visit. No build blocker. | DONE |
| 5 | Does MMI want QC / Calibration data sent to LIMS, or patient results only? | CONFIRMED: Patient results only for now. QC/Calibration to be enabled in future via config flag. | DONE |
| 6 | Serial Number and Firmware version of the specific AU480 at MMI? | CONFIRMED: SN 2023070728 ¦ REF B96693 ¦ Manufactured July 2023, Japan | DONE |
| 7 | Is there a null modem adapter or straight-through RS-232 cable at MMI? | CONFIRMED: Straight-through cable. Standard wiring applies: TxD pin 3 -> RxD pin 2, RxD pin 2 -> TxD pin 3, GND pin 5 -> GND pin 5. | DONE |

# 8. Reusability - Future Labs with AU480

Once MMI integration is complete and live, every future lab that has an AU480 follows this process:

| Step | Action | Time Estimate |
| --- | --- | --- |
| 1 | Copy au480_mmi.json to au480_[newlab].json | 5 minutes |
| 2 | Update connection.port to the COM port at new lab | 2 minutes |
| 3 | Run INSERT into lab_analyzers for new lab | 5 minutes |
| 4 | Map online test numbers to new lab's lims_parameter_uid values | 30-60 minutes |
| 5 | Start integration engine with new config | 5 minutes |
| 6 | Run simulation test | 30 minutes |
| 7 | Go live | Same day |

Third-Party LIS Compatibility: The lis_integration_results table acts as a neutral staging layer. For labs that do not use SpeciGo LIMS, we build an alternate output adapter that reads from lis_integration_results and pushes via webhook, file export, or TCP to the third-party system. The AU480Parser.js and transport layer are fully reused.

# 9. Exact LIMS Integration Points

This section specifies precisely where and how the integration engine touches the existing SpeciGo LIMS codebase. Every table, every field, every query is defined here. No guesswork required.

## 9.1 The Result Writing Query - Exact SQL

When the AU480 sends a result for a sample, the integration engine runs this sequence of queries in a single MySQL transaction:

-- STEP 1: Verify the barcode exists and is active in the system
SELECT rb.barcode_uid, rb.lab_uid, rb.report_uid, rb.patient_uid,
       rb.processing_lab_uid, rb.processing_lab_type, rb.parent_barcode_uid,
       rb.test_type, rb.package_uid, rb.test_uid, rb.sample_type_uid,
       rb.collection_condition, rb.sub_dept_uid
FROM report_barcode rb
WHERE rb.barcode_uid = '[sample_id_from_au480]'
  AND rb.barcode_is_active = 1
LIMIT 1;

      
-- STEP 2: Resolve online_test_no to lims_parameter_uid
SELECT lap.lims_parameter_uid, lap.lims_test_uid, lap.unit
FROM lab_analyzer_parameters lap
WHERE lap.lab_uid = '[lab_uid_from_step1]'
  AND lap.analyzer_code = 'AU480'
  AND lap.online_test_no = '[3-digit-online-test-no]'
  AND lap.is_active = 1
LIMIT 1;

      
-- STEP 3: Write the result value into report_test_values_data
UPDATE report_test_values_data
SET
  value = '[numeric_value_as_string]',
  quantitative_value = '[numeric_value_as_string]',
  param_remark = '[data_flag_meaning]',
  updated_at = NOW()
WHERE barcode_uid = '[barcode_uid]'
  AND parameter_uid = '[lims_parameter_uid_from_step2]'
  AND lab_uid = '[lab_uid]';

      
-- STEP 4: Update test status to TC (Test Completed) in report_tests
-- Only update if ALL parameters for this barcode now have values
UPDATE report_tests
SET test_status = 'TC', updated_at = NOW()
WHERE barcode_uid = '[barcode_uid]'
  AND lab_uid = '[lab_uid]'
  AND test_active_status = 'ACTIVE'
  AND test_class = 'ROUTINE';

      
-- STEP 5: Insert into lis_integration_results for audit trail
INSERT INTO lis_integration_results
  (lab_uid, analyzer_uid, sample_id, online_test_no, parameter_code,
   lims_parameter_uid, raw_value, numeric_value, data_flag,
   result_category, mapping_status, lims_write_status, lims_barcode_uid, received_at)
VALUES
  ('[lab_uid]', '[analyzer_uid]', '[barcode_uid]', '[online_test_no]', '[param_code]',
   '[lims_parameter_uid]', '[raw_6char_string]', [float_value], '[flag]',
   'PATIENT', 'MAPPED', 'WRITTEN', '[barcode_uid]', NOW());

The test_status enum in report_tests is: TP (Test Pending), TS (Test Started), PC (Partial Complete), TC (Test Completed), RS (Resample), RT (Retest), TR (Test Rejected). The integration engine sets TC when a result is successfully written. If only some parameters of a multi-parameter test are received, it sets PC.

## 9.2 Handling the data_flag in param_remark

The AU480 data flag is stored in param_remark as a human-readable string, not the raw code. This ensures lab technicians see meaningful text in the LIMS report view:

| AU480 Flag | param_remark Value Written | Meaning |
| --- | --- | --- |
| N + space | Normal | Within reference range |
| H + space | High | Above reference range |
| L + space | Low | Below reference range |
| ph | Critical High | Above panic value - alert required |
| pl | Critical Low | Below panic value - alert required |
| # | Insufficient Sample | Sample volume inadequate |
| % | Clot Detected | Sample clotted - recollect |
| ? | Calculation Error | Result could not be computed |
| F | Above Range | Result exceeds dynamic range - dilute and rerun |
| G | Below Range | Result below dynamic range |
| R | Reagent Low | Insufficient reagent during run |
|  | NULL | No flag - value not stored in param_remark |

## 9.3 Windows Service Setup at MMI

The integration engine runs as a Windows background service on the lab PC at MMI so it restarts automatically if the PC reboots. The setup uses node-windows npm package:

-- File: install-service.js (run once at MMI during setup)
const Service = require('node-windows').Service;

      
const svc = new Service({
  name: 'SpeciGo LIS Integration Engine',
  description: 'SpeciGo AU480 LIS Integration - MMI Diagnostics',
  script: 'C:\speciго-lis-engine\index.js',
  nodeOptions: [],
  wait: 2,
  grow: 0.5
});

      
svc.on('install', () => { svc.start(); });
svc.install();

      
-- Run once:
node install-service.js

      
-- Service then appears in Windows Services as:
-- 'SpeciGo LIS Integration Engine'
-- Startup type: Automatic
-- Logs written to: C:/speciго-lis-engine/daemon/

## 9.4 REST API Endpoints - Node.js Routes to Add

These are the API endpoints the integration engine exposes. They are called by the React frontend LIS dashboard. Add these routes to your existing Express application:

-- Route file: routes/lis.routes.js

      
GET  /api/lis/status/:lab_uid
  Returns: { connected: bool, port: 'COM2', analyzer: 'AU480',
             lastMessageAt: timestamp, messagesLast24h: int }

      
GET  /api/lis/results/:lab_uid
  Query params: ?date=YYYY-MM-DD&status=WRITTEN|FAILED|PENDING&limit=50
  Returns: Array of lis_integration_results records

      
GET  /api/lis/results/:lab_uid/:sample_id
  Returns: All results for a specific barcode_uid

      
POST /api/lis/mapping/:lab_uid
  Body: { analyzer_code, online_test_no, lims_parameter_uid, lims_test_uid, unit }
  Action: INSERT/UPDATE lab_analyzer_parameters
  Returns: { success: true, mapping_id: int }

      
GET  /api/lis/mapping/:lab_uid
  Returns: All parameter mappings for this lab and analyzer

      
GET  /api/lis/analyzers/:lab_uid
  Returns: All analysers registered for this lab from lab_analyzers

      
POST /api/lis/analyzers/:lab_uid
  Body: Complete analyser record
  Action: INSERT into lab_analyzers with extended fields

      
GET  /api/lis/log/:lab_uid
  Query params: ?limit=100&type=ERROR|CONNECT|DISCONNECT
  Returns: lis_integration_log records

## 9.5 React Frontend - LIS Dashboard Component Location

Add the LIS Dashboard as a new protected route in your existing React application. Based on your LIMS architecture, it sits alongside the existing lab management screens:

-- New route in your React Router config:
<Route path='/lab/lis-integration' element={<LISDashboard />} />

      
-- Component file: src/components/LIS/LISDashboard.jsx
-- Sub-components:
--   src/components/LIS/AnalyserStatus.jsx      <- Live connection status
--   src/components/LIS/ResultsTable.jsx         <- Recent results from lis_integration_results
--   src/components/LIS/ParameterMapping.jsx     <- Map online_test_no to LIMS params
--   src/components/LIS/AnalyserMaster.jsx       <- Multi-analyser library browser
--   src/components/LIS/IntegrationLog.jsx       <- Communication log viewer

      
-- Navigation: Add to sidebar under Lab Management section
-- Icon: Use existing icon pattern from your UI
-- Access control: LABORATORY role only (same as existing lab screens)

The React dashboard communicates with the integration engine via the REST API endpoints defined in section 9.4. The integration engine runs locally on the lab PC on a configurable port (default 3001). Your cloud LIMS backend proxies requests to it. This means the dashboard works even when the lab PC is behind NAT.

# 10. Analyser Master UI - Design Specification

This is the admin interface where SpeciGo manages its library of all analysers across all manufacturers. It has two levels: the Analyser Library (global, across all labs) and the Lab Analyser Config (per-lab instance of a known analyser).

## 10.1 Screen 1 - Analyser Library (Global)

This screen is accessible to SpeciGo admin only. It shows all analysers that SpeciGo has ever integrated, regardless of which lab they are at. Think of it as your catalogue.

| UI Element | Specification |
| --- | --- |
| Page Title | Analyser Integration Library |
| Filter Bar | Filter by: Manufacturer / Category / Protocol / Integration Status |
| Search | Search by model name or manufacturer |
| Analyser Card | Shows: Logo, Manufacturer, Model, Category, Protocol, No. of Labs Using, Status badge (Live/Testing/Not Started) |
| Add New Button | Opens Add Analyser form - captures all fields from Section 2 of this doc |
| Click Card | Opens Analyser Detail page (Screen 2) |
| Export | Download analyser library as CSV/PDF for documentation |

## 10.2 Screen 2 - Analyser Detail Page

This screen shows everything about one analyser model. It has tabs:

| Tab | Content |
| --- | --- |
| Identity | All fields from Section 2.1: Manufacturer, Model, Category, Throughput, Sample Types, etc. |
| Communication | All fields from Section 2.2: Protocol, Baud Rate, Parity, Class, Start/End codes etc. |
| Parameters | Full parameter list with online test numbers, LIMS names, units, LOINC codes |
| Labs Using | List of all labs that have this analyser, with their integration status and last active timestamp |
| Notes | Known quirks, edge cases, cable requirements, firmware notes - free text |
| Driver Files | Reference to which Node.js config and parser files handle this analyser |

## 10.3 Screen 3 - Lab Analyser Config (Per Lab)

This screen is accessible to lab admin. It shows the analysers configured for their specific lab and allows mapping of parameters to their LIMS test master.

| UI Element | Specification |
| --- | --- |
| Analyser List | All analysers registered for this lab from lab_analyzers table |
| Status Indicator | Live green dot / Red disconnected / Yellow testing - pulled from /api/lis/status |
| Connection Details | COM port, Baud rate, Protocol - read-only display |
| Parameter Mapping Table | Left: Online Test No + Analyser Name ¦ Right: Dropdown to select LIMS test parameter |
| Save Mapping Button | Calls POST /api/lis/mapping/:lab_uid to save all mappings |
| Recent Results | Last 20 results received from this analyser with status indicators |
| Integration Log | Last 50 communication events - connect, disconnect, errors, message counts |
| Test Connection | Button that sends a test query and shows if COM port responds |

## 10.4 Screen 4 - Live Results Monitor

This is the real-time view that lab technicians use during a run. It auto-refreshes every 5 seconds:

| UI Element | Specification |
| --- | --- |
| Header Strip | Analyser name, COM port, Connected status, Messages received today |
| Results Feed | Scrolling table: Time ¦ Barcode UID ¦ Patient Name ¦ Test ¦ Result ¦ Flag ¦ Write Status |
| Flag Colour Coding | Normal=green, High/Low=amber, Critical=red, Error flags=dark red |
| Write Status | WRITTEN=green tick, PENDING=clock icon, FAILED=red cross with error message on hover |
| Filter | Filter by date, sample ID, write status, flag type |
| Click Row | Opens the full patient report in your existing report view |
| Export | Export results for date range as CSV |

# 11. Multi-Manufacturer Analyser Framework

This section defines how SpeciGo expands beyond the AU480 to build a library of analysers across manufacturers. Every future analyser follows this exact framework so the process is repeatable and the codebase stays clean.

## 11.1 Manufacturer Categories and Protocol Families

Most analysers in Indian labs fall into one of four protocol families. Once you have built a driver for one analyser in a family, the parser is largely reusable for others in the same family:

| Protocol Family | Manufacturers | Notes |
| --- | --- | --- |
| Beckman Coulter Proprietary | Beckman Coulter AU series, DxH series | Our AU480 driver. AU680 reuses 90% of it. |
| Standard ASTM E1394 | Sysmex, Mindray, Horiba, Erba | Generic ASTM parser covers most of these |
| HL7 v2.x | Roche, Abbott, Siemens high-end | HL7 parser needed, LLP framing instead of STX/ETX |
| Proprietary Binary/CSV | Various POC devices, some Agappe | Fully custom per-analyser, spec sheet mandatory |

## 11.2 Information Collected Per New Analyser

For every new analyser that SpeciGo integrates, this is the mandatory information collection checklist. Claude will extract and structure this from the spec sheet provided:

| Information Item | Source | Mandatory? |
| --- | --- | --- |
| Manufacturer + Model + REF | Machine label (photo) | Yes |
| Serial Number | Machine label (photo) | Yes |
| Manufacture Date | Machine label (photo) | Yes |
| Protocol family (ASTM/HL7/proprietary) | Spec sheet | Yes |
| Communication type (serial/TCP/LAN) | Spec sheet + site visit | Yes |
| Baud rate, data bits, parity, stop bits | Analyser screen photo | Yes (serial only) |
| TCP port and connection mode | Analyser screen photo | Yes (TCP only) |
| Start/end codes + framing details | Spec sheet | Yes |
| Message format for result records | Spec sheet - full structure | Yes |
| Full parameter list with test codes/numbers | Analyser screen photo | Yes |
| Error/flag code list | Spec sheet appendix | Yes |
| Cable type at site (null modem/straight) | Site visit | Yes (serial only) |
| COM port at site Windows PC | Site visit | Yes (serial only) |
| IP address and port at site | Site visit | Yes (TCP only) |
| Barcode format | LIMS system | Yes |
| Interface direction (uni/bi) | Lab requirement | Yes |
| QC/Calibration data required? | Lab requirement | Yes |
| ISE module present? | Site visit | If applicable |
| Firmware version | Analyser screen | Recommended |

## 11.3 Per-Analyser File Structure

Every new analyser gets exactly these files. Nothing more, nothing less:

speciго-lis-engine/
├── config/analysers/
│   ├── au480_mmi.json              <- AU480 at MMI (BUILT - v1.5)
│   ├── au480_[labname].json        <- AU480 at future lab (copy + update port)
│   ├── sysmex_xn550_[lab].json     <- Future: Sysmex XN-550 (ASTM)
│   └── mindray_bc6800_[lab].json   <- Future: Mindray BC-6800 (ASTM)
├── src/protocol/
│   ├── AU480Parser.js              <- Beckman Coulter family parser (BUILT)
│   ├── ASTMParser.js               <- Reusable for Sysmex, Mindray, Horiba
│   ├── HL7Parser.js                <- Future: Roche, Abbott, Siemens
│   └── [Manufacturer]Parser.js     <- Only if truly proprietary
├── src/transport/
│   ├── SerialPortManager.js        <- Reused by ALL serial analysers (BUILT)
│   └── TCPSocketManager.js         <- Reused by ALL TCP analysers (BUILT)

## 11.4 Adding a New Analyser - The Standard Process

This is the repeatable process for every future analyser. It takes one Claude Project conversation per new analyser:

| Step | Action | Who Does It | Time |
| --- | --- | --- | --- |
| 1 | Share spec sheet PDF and machine label photo with Claude | Gagandeep | 5 mins |
| 2 | Share analyser screen photos (protocol settings, test numbers) | Gagandeep / Umesh at site | 15 mins |
| 3 | Answer 7-question checklist (same as AU480 process) | Gagandeep | 15 mins |
| 4 | Claude generates foundation document for new analyser | Claude | Auto |
| 5 | Claude generates config JSON file | Claude | Auto |
| 6 | Claude generates or adapts parser (reuse if same protocol family) | Claude | Auto |
| 7 | Claude generates test simulation script for new analyser | Claude | Auto |
| 8 | Umesh deploys at site, maps parameters via UI | Umesh | 1-2 hrs |
| 9 | Site test and go live | Umesh | 1-2 hrs |

## 11.5 Priority Analyser Roadmap

Based on common analysers found in Indian NABL labs, here is the recommended order to build your library after the AU480:

| Priority | Analyser | Category | Protocol | Parser Needed |
| --- | --- | --- | --- | --- |
| 1 | Beckman Coulter AU480 (MMI) | Biochemistry | BC Proprietary | AU480Parser.js - BUILDING NOW |
| 2 | Beckman Coulter AU680 | Biochemistry | BC Proprietary | AU480Parser.js reused (95%) |
| 3 | Sysmex XN-550 / XN-1000 | Haematology | ASTM E1394 | ASTMParser.js (new) |
| 4 | Mindray BC-6800 | Haematology | ASTM E1394 | ASTMParser.js reused |
| 5 | Sysmex CS-2500 / CA-560 | Coagulation | ASTM E1394 | ASTMParser.js reused |
| 6 | Roche Cobas c311 | Biochemistry | HL7 v2.x | HL7Parser.js (new) |
| 7 | Agappe Mispa i2 / i3 | Immunoassay | Proprietary | AgappeParser.js (new) |
| 8 | Erba XL-200 / XL-640 | Biochemistry | ASTM E1394 | ASTMParser.js reused |
| 9 | Horiba Yumizen H500 | Haematology | ASTM E1394 | ASTMParser.js reused |
| 10 | Siemens Atellica | Multi-discipline | HL7 v2.x | HL7Parser.js reused |

By priority 3, you will have both the BC Proprietary parser and the ASTM parser built. From priority 4 onwards, adding a new analyser primarily requires only a new config JSON file and a parameter mapping exercise. The code is already written. This is the compounding value of the analyser library.

END OF DOCUMENT - v1.5

Upload this document to the SpeciGo LIS Integration Claude Project as the foundation context.